Galvant Industries GPIBUSB Command List

** Note: All writes to the controller must end with a carriage return ('\r' or dec:13) **
** All responses from the adapter will have a CR added. **

+a:1
Set the GPIB address of the device you wish to communicate with. Here the target address is 1.

+eoi:1
Used to toggle EOI detection on (1) and off (0). If set to on, the controller will stop reading bytes
from the device when EOI is asserted. When set to off, the controller will stop when it reads the byte
defined by the +eos: command. Default is on (1).

+eos:1
Set the character that the device uses to signal end of data. This is especially important when EOI
is off. However, some instruments will use both a termination character and use EOI. It is recommended that you disable any termination characters at the instrument and just use EOI.
The controller will terminate transmissions with attached computer with a carrage return ('\r' or dec:13).
This extra byte is NOT sent IF +eos:13 is used. This prevents double sending of carrage return bytes.
To set the termination character, put the decimal form of the character after the colon. For example, to
set the character to 'z', you would write +eos:122

+read
Force the controller to read data from the instrument. Normally this is NOT required. When writing
commands that end with a question mark (?), the controller will automatically read the response.
However, older instruments do not include question marks at the end of query commands. In these cases, using this command will prompt the controller to read the response data.

+test
Used to test the controller and your communication software. Controller will respond with: testing\n\r
Controller needs to be plugged into an instrument that is on for this to work.